## README
RTMify — FMEA Template
Failure Mode and Effects Analysis for Regulated Products · RTMify.io
Standards coverage: |  | ISO 14971 | ISO 13485 | IEC 62304 | AS9100 | ISO 26262 | DO-178C
WHAT THIS IS
A working Failure Mode and Effects Analysis (FMEA) spreadsheet for regulated product development. Three tabs: this README, the FMEA worksheet, and a Scales reference. Covers the risk analysis and evaluation required by ISO 14971 (medical device risk management), ISO 13485 §7.1 (risk-based decision making), IEC 62304 §7 (software risk management), AS9100 §6.1 (actions to address risks and opportunities), ISO 26262 Part 5 (hardware FMEA), and DO-178C §2.4 (safety assessment). Use it standalone or alongside the RTMify Requirements Tracking Template — the Linked REQ and Linked Risk ID columns cross-reference directly.
THE TABS
Tab | What goes here
FMEA | One row per failure mode. Identify the item or function, how it could fail, what happens if it does, and how likely and detectable the failure is. Score Severity, Occurrence, and Detection on 1–5 scales (matching the RTM Risks tab). RPN = S × O × D is calculated automatically. After scoring: document recommended actions, assign responsibility, and re-score after mitigation. The Linked REQ and Linked Risk ID columns let you cross-reference this FMEA to the RTM template. Every risk control measure should trace to a requirement, and every FMEA finding should map to a risk register entry.
Scales | Reference tables for Severity, Occurrence, and Detection ratings. Each scale runs 1–5 with descriptions specific to medical device and regulated product contexts. Adapt the descriptions to your product and risk management plan. The scales should reflect your organization's risk acceptance criteria, not a generic template.
FIELD REFERENCE — FMEA TAB
FMEA ID | Unique identifier for each failure mode entry (FMEA-001, FMEA-002…). One failure mode per row. If an item has multiple failure modes, each gets its own row and ID.
Item / Function | The component, subsystem, or function being analyzed. Be specific: "Temperature sensor analog front-end" not "sensor."
Potential Failure Mode | How this item or function could fail to perform its intended purpose. State the failure, not the cause: "Reads 2°C high" not "Solder joint cracks."
Potential Effect(s) | Downstream consequence of the failure mode, including harm to patient, user, or operator. ISO 14971 requires you to identify the hazardous situation and the resulting harm.
Severity (S) | How serious is the effect if the failure occurs? 1 = negligible, 5 = catastrophic. See Scales tab. Severity almost never changes with mitigation. If it does, document why.
Potential Cause(s) | Root cause or mechanism that produces the failure mode. One failure mode may have multiple causes; use separate rows or comma-separate if causes share the same S/O/D profile.
Occurrence (O) | How likely is this cause to produce the failure mode? 1 = remote, 5 = frequent. Based on historical data, test results, or engineering judgment.
Current Controls — Prevention | Existing design or process controls that prevent the cause or failure mode from occurring. "TCXO specified in BOM" or "Input range check in firmware."
Current Controls — Detection | Existing controls that detect the failure mode before it reaches the user. "Power-on self-test" or "Automated boundary scan during manufacturing."
Detection (D) | How likely are current controls to detect the failure before it reaches the user? 1 = certain detection, 5 = undetectable. Note: scale is inverted — lower is better.
RPN | Risk Priority Number = S × O × D. Calculated automatically. Range 1–125 with 1–5 scales. Use RPN to prioritize actions, not as a pass/fail threshold.
Recommended Action | What should be done to reduce the risk? Target the highest contributor: reduce Occurrence (better design), improve Detection (better testing), or both. Severity reduction requires a fundamental design change.
Responsible / Target Date | Who owns the action and when is it due? These columns exist so the FMEA is a living document, not a one-time exercise.
Action Taken | What was actually done? Document the implemented mitigation, not just the plan.
Revised S / O / D / RPN | Re-score after mitigation is implemented. Revised RPN is calculated automatically. If residual risk is still unacceptable, iterate.
Linked REQ | Requirement ID from the RTMify Requirements Tracking Template. Every risk control measure should trace to a verifiable requirement.
Linked Risk ID | Risk ID from the RTM Risks tab. Maps this FMEA entry to the risk register for bidirectional traceability between detailed FMEA analysis and the summary risk assessment.
Notes | Rationale, references, audit notes, standard-specific annotations. Use freely.
STANDARD-SPECIFIC NOTES
ISO 14971:2019 | This is the primary standard governing risk management for medical devices. The FMEA is one method (among fault tree analysis, hazard analysis, etc.) for risk analysis per §5.4. ISO 14971 requires: identification of hazards and hazardous situations (Failure Mode + Effect columns), estimation of risk (Severity × Occurrence), evaluation against risk acceptance criteria (your risk management plan defines acceptable RPN thresholds), risk control (Recommended Action → Action Taken), and verification of risk control effectiveness (Revised scores). Residual risk must be evaluated. The Revised RPN column serves this purpose. The Linked REQ column satisfies §5.5: each risk control measure must be traceable to a verifiable requirement.
ISO 13485:2016 | §7.1 requires risk-based decision making throughout the product lifecycle. §7.3.3 (design inputs) must include risk management outputs — FMEA findings feed design requirements. §7.3.4 (design outputs) must reference risk analysis results. This FMEA template, used alongside the RTM template, provides the traceability chain auditors look for: FMEA finding → risk control requirement → verification test.
IEC 62304:2006+A1 | §7 requires software risk management integrated with ISO 14971. For software-specific failure modes (race conditions, buffer overflows, state machine errors), this FMEA captures the analysis. Software safety classification (A/B/C) drives which failure modes require analysis. Class C software requires analysis of all failure modes; Class A may use a lighter approach. Note the software safety class in the Notes column.
AS9100 Rev D | §6.1 requires the organization to determine risks and opportunities that need to be addressed. While AS9100 does not mandate FMEA specifically, it is the most common method aerospace organizations use to satisfy §6.1 for product-level risk analysis. The FMEA also feeds §8.3.5 design verification. Failure modes identified here should have corresponding verification activities in the test plan.
ISO 26262:2018 | Part 5 §7 covers hardware FMEA. Part 11 describes the FMEA method in detail. ASIL decomposition from the HARA drives which components require FMEA — ASIL C and D typically require systematic FMEA for all safety-related hardware. Note the ASIL in the Notes column for each entry. For software, fault tree analysis (Part 6 §7) is more common than FMEA, but both are accepted methods.
USING WITH THE RTM TEMPLATE
This FMEA template is designed to work alongside the RTMify Requirements Tracking Template. The connection points: • Linked REQ column → references requirement IDs (REQ-001, REQ-002…) from the RTM Requirements tab • Linked Risk ID column → references risk IDs (RSK-101, RSK-102…) from the RTM Risks tab • Severity and Occurrence use the same 1–5 scales as the RTM Risks tab for consistency • Every Recommended Action that becomes a risk control should be captured as a requirement in the RTM Download the RTM template at RTMify.io/download if you don't already have it.

## FMEA
FMEA ID | Item / Function | Potential Failure Mode | Potential Effect(s) | Severity (S) | Potential Cause(s) | Occurrence (O) | Current Controls — Prevention | Current Controls — Detection | Detection (D) | RPN | Recommended Action | Responsible | Target Date | Action Taken | Revised S | Revised O | Revised D | Revised RPN | Linked REQ | Linked Risk ID | Notes
FMEA-001 | Temperature sensor analog front-end | Reads 2°C high under sustained load | Incorrect temperature compensation applied; dosage calculation error; potential patient harm from over/under-dosing | 4 | Thermal self-heating of sensor element at high sampling rates | 3 | Thermal simulation during design phase | Power-on self-test calibration check | 2 | 24 | Add external TCXO; reduce sampling rate during sustained operation | HW Lead | 2026-04-15 | TCXO added per BOM rev C; sampling rate governed by firmware | 4 | 1 | 2 | 8 | REQ-602 | RSK-101 | ISO 14971 §5.4 — thermal drift is a known hazard for this sensor family
FMEA-002 | Software watchdog timer | Watchdog fails to trigger on main loop hang | System unresponsive; alarm notifications delayed; clinician unaware of critical patient state change | 5 | Watchdog kick occurs in ISR independent of main loop health | 2 | Code review checklist item for watchdog architecture | Integration test: inject main loop stall, verify watchdog reset | 2 | 20 | Refactor watchdog kick to main loop; add secondary watchdog in safety co-processor | FW Lead | 2026-05-01 | [6 empty] | RSK-205 | IEC 62304 Class C — requires analysis of all software failure modes
FMEA-003 | Vital signs display rendering | Stale data displayed after communication loss with sensor module | Clinician sees last-known-good values instead of current state; delayed intervention; potential patient harm | 4 | Display refresh timer continues without checking data freshness timestamp | 2 | Requirements specify data staleness timeout (REQ-415) | System-level validation test: simulate comm loss, verify timeout indicator | 3 | 24 | Add staleness indicator overlay when data age exceeds 3 seconds; grey out affected readings | SW Lead | 2026-04-30 | [5 empty] | REQ-415 | RSK-310
FMEA-004 | Rechargeable battery pack | Capacity degrades below minimum operating threshold within rated cycle life | Unexpected shutdown during patient monitoring; loss of data and alarm capability | 4 | Cell aging accelerated by sustained high-temperature charging | 2 | Charge management IC limits charge rate above 40°C | Battery health check at power-on; capacity estimation via coulomb counting | 2 | 16 | Add predictive battery replacement warning at 80% rated capacity; log charge temperature histogram | HW Lead | 2026-06-01

## Scales
SEVERITY (S) — How serious is the effect if the failure occurs?
Rating | Level | Description (adapt to your product and risk management plan)
1 | Negligible | No injury or health impact. Nuisance-level effect. No regulatory reporting required.
2 | Minor | Temporary discomfort or minor inconvenience. No medical intervention needed. Reversible without treatment.
3 | Moderate | Injury or impairment requiring medical intervention. Temporary but clinically significant. May require adverse event reporting.
4 | Serious | Serious injury, permanent impairment, or life-threatening situation. Requires immediate medical intervention. Reportable event.
5 | Catastrophic | Death or irreversible serious harm. Recall-level severity.
OCCURRENCE (O) — How likely is this cause to produce the failure?
Rating | Level | Description (adapt to your product and risk management plan)
1 | Remote | Failure is unlikely based on field data and engineering analysis. < 1 in 100,000 device-uses.
2 | Low | Isolated failures observed in similar devices or test conditions. 1 in 10,000 to 1 in 100,000.
3 | Moderate | Occasional failures expected based on design complexity or known failure mechanisms. 1 in 1,000 to 1 in 10,000.
4 | High | Repeated failures likely based on similar prior designs or known weaknesses. 1 in 100 to 1 in 1,000.
5 | Frequent | Failure is almost certain without intervention. > 1 in 100 device-uses.
DETECTION (D) — How likely are current controls to catch the failure? (lower = better)
Rating | Level | Description (adapt to your product and risk management plan)
1 | Certain | Current controls will almost certainly detect the failure mode. Automated 100% inspection or built-in self-test with proven reliability.
2 | High | Controls detect most occurrences. Automated testing with good fault coverage, or rigorous inspection protocols.
3 | Moderate | Controls detect some occurrences. Sample-based testing, manual inspection, or design review. Some failure modes may escape.
4 | Low | Controls unlikely to detect the failure mode. Failure depends on field observation, customer complaint, or post-market surveillance.
5 | Undetectable | No current controls exist to detect the failure mode before it reaches the user. Discovery is accidental or by adverse event report.
RPN GUIDANCE
RPN = Severity × Occurrence × Detection. Range is 1–125 with 1–5 scales. RPN thresholds are guidance, not absolute rules. Your risk management plan should define acceptable risk levels based on your product's risk-benefit profile and applicable regulatory requirements. A high-severity, low-occurrence, low-detection item (e.g., S=5, O=1, D=1 → RPN=5) may still require explicit risk acceptance documentation even though its RPN is low. Common thresholds (adapt to your risk management plan): • RPN > 40: Action required — unacceptable risk without additional mitigation • RPN 15–40: Action recommended — evaluate cost-benefit of additional controls • RPN < 15: Monitor — document rationale for risk acceptance ISO 14971 §5.5 requires that risk evaluation use criteria established in the risk management plan, not arbitrary RPN cutoffs. The RPN is a prioritization tool, not a compliance threshold.